## System Configuration
Scenario 7.1: NEDT Reconciliation — System Configuration
As-built MWIR sensor, lab/TVAC test configuration
Optics
Primary aperture diameter | 30 | cm | As-built, measured on optical bench
Effective focal length | 121.5 | cm | Measured via star test (nominal 120 cm)
f-number (actual) | 4.05 | — | Derived from D and f (nominal f/4.0)
Optical transmission | 71 | % | Measured at λ = 4.25 µm, includes all elements
Optics temperature | 22 | °C | Lab ambient (TVAC shroud at 22°C)
Number of optical elements | 4 | — | Primary, secondary, fold, window
Spectral Band
Filter cut-on | 3500 | nm | Cold filter, measured at 77 K
Filter cut-off | 5000 | nm | Cold filter, measured at 77 K
Band center | 4250 | nm | Arithmetic center
Test Setup
Blackbody emissivity | 0.995 | — | CI Systems SR-800R extended area blackbody
Blackbody fill | overfills FOV | — | 30×30 cm aperture at 50 cm working distance
Chamber pressure | <1e-5 | Torr | TVAC chamber, turbomolecular pump
Shroud temperature | 22 | °C | Ambient lab temperature
Shroud emissivity | 0.95 | — | Anodized aluminum shroud
Cold stop efficiency | 1 | — | Designed; assumed 100% for this test

## As-Built Detector
As-Built Detector Parameters (Measured)
HgCdTe MWIR FPA, 640×512, Lot W2024-037
Parameter | Value | Unit | Notes
Pixel pitch | 18 | µm | Square pixels
Quantum efficiency (in-band avg) | 68 | % | Measured at 77 K, 3.5–5.0 µm average
Dark current (measured) | 135 | e⁻/s | At 77.0 K operating temperature
Operating temperature | 77 | K | LN2 cold finger, ±0.1 K stability
Read noise (post-CDS) | 14.2 | e⁻ RMS | Measured in dark, 200-frame ensemble
Full well capacity | 500000 | e⁻ | 90% linearity limit
ADC resolution | 14 | bits | 14-bit ADC, 2's complement
System gain | 12 | e⁻/DN | Set for 70% FWC at ADC max
Integration time | 0.5 | ms | Lab NEDT test (shorter than 8 ms orbital nominal to avoid well saturation)
IPC coupling | 1.5 | % | Measured via single-pixel reset method
Flicker noise (1/f) | 0 | e⁻ | Below measurement floor at this frame rate
ROIC glow | 5 | e⁻/s | Measured with cold shutter closed
Nominal vs. As-Built
Parameter | Nominal | As-Built | Delta
QE [%] | 72 | 68 | -4
Dark current [e⁻/s] | 100 | 135 | 35
Read noise [e⁻ RMS] | 12 | 14.2 | 2.2
f-number | 4 | 4.05 | 0.05
Transmission [%] | 73 | 71 | -2

## NEDT Measurements
NEDT Measurements — TVAC Test Campaign
Date: 2026-03-28, Operator: K. Martinez, Chamber: TVAC-3
Blackbody T [°C] | Blackbody T [K] | Measured NEDT [mK] | Std Dev [mK] | N frames | Notes
15 | 288.15 | 160 | 8.5 | 500 | Low-temperature point
20 | 293.15 | 142 | 7.2 | 500
25 | 298.15 | 127 | 6 | 500 | Primary test point
30 | 303.15 | 114 | 5.1 | 500
35 | 308.15 | 104 | 4.5 | 500
40 | 313.15 | 95 | 3.8 | 500
50 | 323.15 | 81 | 3 | 500 | High-temperature point
Measurement Method:
NEDT computed as σ(frame-to-frame pixel value) / (dS/dT) from 2-point calibration at ±2°C around each test temperature.
Spatial average of NEDT across central 100×100 pixel ROI (excludes edge pixels and known bad pixels).