## NEDT Comparison
Scenario 7.1: Predicted vs. Measured NEDT
BB T [°C] | BB T [K] | Measured NEDT [mK] | Predicted NEDT [mK] | Δ [mK] | Signal [e⁻] | dS/dT [e⁻/K] | σ_total [e⁻] | SNR
15 | 288.15 | 160 | 138.7 | 21.3 | 97857 | 3758.9 | 521.36 | 187.7
20 | 293.15 | 142 | 123.04 | 18.96 | 118204 | 4392.9 | 540.52 | 218.7
25 | 298.15 | 127 | 110.08 | 16.92 | 141916 | 5105.7 | 562.03 | 252.5
30 | 303.15 | 114 | 99.27 | 14.73 | 169401 | 5903 | 585.97 | 289.1
35 | 308.15 | 104 | 90.18 | 13.82 | 201096 | 6790.9 | 612.42 | 328.4
40 | 313.15 | 95 | 82.5 | 12.5 | 237468 | 7775.2 | 641.43 | 370.2
50 | 323.15 | 81 | 70.33 | 10.67 | 326266 | 10056.9 | 707.27 | 461.3

## Noise Breakdown
NEDT Breakdown by Noise Term (25°C / 298.15 K)
Noise Term | σ [e⁻ RMS] | NEDT_i [mK] | Fraction [%]
signal_shot | 376.72 | 73.784 | 44.9
background_shot | 348.58 | 68.272 | 38.5
nearfield_shot | 228.56 | 44.766 | 16.5
read_noise | 14.2 | 2.781 | 0.1
quantization | 3.46 | 0.678 | 0
dark_shot | 0.26 | 0.051 | 0
straylight_shot | 0 | 0 | 0
gr_noise | 0 | 0 | 0
johnson_noise | 0 | 0 | 0
flicker_1f | 0 | 0 | 0
ktc_reset | 0 | 0 | 0
prnu | 0 | 0 | 0
dsnu | 0 | 0 | 0
clutter | 0 | 0 | 0
persistence_noise | 0 | 0 | 0
glow_shot | 0 | 0 | 0

## Sensitivity
NEDT Sensitivity to Input Parameters
Parameter | Base Value | Unit | NEDT- [mK] | NEDT+ [mK] | Δ/1% [mK]
QE | 0.68 | fraction | 110.634 | 109.532 | -0.5508
Dark current | 135 | e⁻/s | 110.079 | 110.079 | 0
Read noise | 14.2 | e⁻ RMS | 110.078 | 110.08 | 0.0007
Integration time | 0.0005 | s | 110.634 | 109.532 | -0.5508
Optical transmission | 0.71 | fraction | 110.952 | 109.223 | -0.8647
f-number (via focal length) | 1.215 | m | 108.977 | 111.18 | 1.1015

## Summary
Scenario 7.1: NEDT Reconciliation Summary
Test Conditions
Blackbody temperature | 298.15 K (25°C)
Blackbody emissivity | 0.995
Atmosphere | exo (vacuum / TVAC)
Detector temperature | 77.0 K
Integration time | 0.5 ms
Results
Measured NEDT [mK] | 127.0
Predicted NEDT (as-built) [mK] | 110.08
Predicted NEDT (nominal) [mK] | 103.37
Gap (measured − predicted) [mK] | 16.92
Signal at 298.15 K [e⁻] | 141,916
dS/dT [e⁻/K] | 5105.7
Total noise [e⁻ RMS] | 562.03
SNR | 252.5
Missing noise σ [e⁻ RMS] | 323.38
Dominant Noise Term | signal_shot
Most Sensitive Parameter | f-number (via focal length)