## System Configuration
Scenario 7.3: MTF Measurement vs. Prediction
As-built VNIR pushbroom sensor — lab slanted-edge MTF test
Optics
Aperture diameter | 20 | cm | As-built, measured on optical bench
Effective focal length | 60 | cm | Measured via collimator (nominal 60 cm)
f-number | 3 | — | Derived from D and f
Optical transmission | 82 | % | Measured at 650 nm, includes all elements
Optics temperature | 22 | °C | Lab ambient
Central obscuration | 25 | % | Secondary mirror (Cassegrain)
Number of elements | 3 | — | Primary, secondary, corrector
Spectral Band
Filter min | 550 | nm | VNIR band
Filter max | 750 | nm | VNIR band
MTF test wavelength | 650 | nm | Quasi-monochromatic collimator source
Detector
Pixel pitch | 10 | µm | Square pixels, silicon CCD
Fill factor | 100 | % | Full-frame CCD
Quantum efficiency | 85 | % | Measured at 650 nm
Dark current | 50 | e⁻/s | At 263 K operating temperature
Operating temperature | 263 | K | TEC-cooled
Read noise | 8 | e⁻ RMS | CDS readout
Full well capacity | 120000 | e⁻ | 90% linearity
ADC bits | 14 | —
System gain | 8 | e⁻/DN
IPC coupling | 1 | % | Measured via single-pixel illumination
Test Setup
Collimator focal length | 3 | m | Off-axis parabola, λ/20 surface
Target type | slanted edge | — | 4° tilt angle, knife-edge on backlit aperture
Integration time | 2 | ms | Lab MTF test frame rate
Number of frames averaged | 100 | — | Reduces noise in ESF measurement

## Measured MTF
Slanted-Edge MTF Measurement — Cross-Track Axis
Method: ISO 12233 slanted-edge, 4° tilt, 100-frame average
Date: 2026-04-10, Operator: K. Martinez, Setup: Optical bench #2
spatial_frequency_cy_mm | MTF_measured
0 | 1
2.041 | 0.9783
4.082 | 0.997
6.122 | 0.9897
8.163 | 0.9342
10.204 | 0.93
12.245 | 0.9357
14.286 | 0.9115
16.327 | 0.8969
18.367 | 0.8637
20.408 | 0.8677
22.449 | 0.8428
24.49 | 0.8076
26.531 | 0.7979
28.571 | 0.7615
30.612 | 0.7142
32.653 | 0.7046
34.694 | 0.656
36.735 | 0.6545
38.776 | 0.6112
40.816 | 0.5796
42.857 | 0.5425
44.898 | 0.5414
46.939 | 0.4913
48.98 | 0.458
51.02 | 0.4303
53.061 | 0.4152
55.102 | 0.3849
57.143 | 0.3584
59.184 | 0.3322
61.224 | 0.3322
63.265 | 0.2691
65.306 | 0.2436
67.347 | 0.216
69.388 | 0.2154
71.429 | 0.2021
73.469 | 0.1635
75.51 | 0.1337
77.551 | 0.116
79.592 | 0.1214
81.633 | 0.107
83.673 | 0.0894
85.714 | 0.0577
87.755 | 0.0586
89.796 | 0.0453
91.837 | 0.0363
93.878 | 0.0365
95.918 | 0.0181
97.959 | 0.0171
100 | 0.001
Notes:
Frequency is in cycles/mm on the focal plane.
MTF normalized to 1.0 at DC (zero frequency).
Nyquist frequency = 50 cy/mm (for 10 um pixels).
Data extends to 2× Nyquist (100 cy/mm) for aliasing analysis.

## As-Built WFE
As-Built Wavefront Error — Interferometric Measurement
4D Technology PhaseCam 6110, 633 nm HeNe
Parameter | Value | Unit | Notes
Total WFE RMS | 0.07 | waves | At 633 nm HeNe reference
WFE reference wavelength | 633 | nm | HeNe laser
WFE P-V | 0.35 | waves | Peak-to-valley at 633 nm
Dominant aberration | trefoil | — | Likely from primary mirror mount stress
Measurement uncertainty (2σ) | 0.005 | waves | Repeatability over 10 measurements
Measurement temperature | 22 | °C | Lab ambient during interferometric test
Measurement date | 2026-04-08 | —

## Focus Position
Focus Position Measurement
Through-focus MTF scan at 650 nm
Parameter | Value | Unit | Notes
Defocus from best focus | 5 | µm | Measured via through-focus scan
Best focus position | 600.023 | mm | Micrometer reading at peak MTF
Current focus position | 600.028 | mm | Operating position (5 µm beyond best)
Focus measurement method | through-focus MTF | — | Peak of MTF@Nyquist vs. focus curve
Focus step size | 1 | µm | Micrometer resolution
Focus uncertainty (2σ) | 2 | µm | Estimated from repeated measurements
Defocus Sweep
Defocus [µm]
0
1
2
3
5
7
10
15
20